## Sheet1
analysis_date | analysis_batch | location | name | measurement_time | pre_acid_time | post_acid_tiem | extra_time | sample_v | rinsing_v | filling_speed | rinsing_speed | pumping_speed | acid_strokes | nb_of_rinses | rinsing_mode | rinse_for_replicates | database_name | filename | flag | salinity | DIC | Area | CV% | Notes
29/09/2021 | 1 | J | J01 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J01.dat | 7 | 32 | -9999 | -9999 | -9999 | Aborted - Kloehn not connecting to laptop
29/09/2021 | 1 | J | J02 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J02.dat | 7 | 32 | 306.6 | 22830 | 0.61
29/09/2021 | 1 | J | J03 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J03.dat | 7 | 32 | 308.4 | 22966 | 0.113 | Weird calculations - restarting system but NOT "reset zero" for Li-cor (restarted after J03)
29/09/2021 | 1 | J | J04 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J04.dat | 7 | 32 | 308.6 | 22977 | 0.214
29/09/2021 | 1 | J | J05 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J05.dat | 7 | 32 | 3270.1 | 24347 | 0.08 | RESET ZERO before J05 + changed CALCULATION MODE to "reject first" (also before running J05)
29/09/2021 | 1 | J | J06 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J06.dat | 7 | 32 | 3270.8 | 24353 | 0.051
29/09/2021 | 1 | J | J07 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J07.dat | 7 | 32 | -9999 | -9999 | -9999 | Aborted - pressed wrong button
29/09/2021 | 1 | CRM | CRM-189-0099-1000 | 360 | 75 | 120 | 10 | 1000 | 1000 | 150 | 150 | 15 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-189-0099-1000.dat | 2 | 33.494 | 3033.2 | 18819 | 0.011
29/09/2021 | 1 | CRM | CRM-189-0099-1200 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-189-0099-1200.dat | 2 | 33.494 | 3041.8 | 22647 | 0.06
29/09/2021 | 1 | CRM | CRM-189-0099-1400 | 360 | 75 | 120 | 10 | 1400 | 1400 | 150 | 150 | 21 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-189-0099-1400.dat | 2 | 33.494 | 3038.3 | 26391 | 0.243
29/09/2021 | 1 | R | R13 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)R13.dat | 2 | 32 | 3277.8 | 24404 | 0.053 | First area = a bit of a vaccum
29/09/2021 | 1 | R | R20 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)R20.dat | 2 | 32 | 3277.2 | 24399 | 0.018
29/09/2021 | 1 | U | U14 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)U14.dat | 2 | 32 | 3281.3 | 24428 | 0.05
29/09/2021 | 1 | U | U42 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)U42.dat | 2 | 32 | 3277.9 | 24402 | 0.098
29/09/2021 | 1 | R | R12 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)R12.dat | 2 | 32 | 3280.2 | 24420 | 0.066
29/09/2021 | 1 | U | U49 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)U49.dat | 2 | 32 | 3274.2 | 24374 | 0.055
29/09/2021 | 1 | U | U06 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)U06.dat | 2 | 32 | 3270.5 | 24350 | 0.096
29/09/2021 | 1 | R | R17 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)R17.dat | 2 | 32 | 3271.3 | 24553 | 0.043
29/09/2021 | 1 | R | R08 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)R08.dat | 2 | 32 | 3272.7 | 24363 | 0.116
29/09/2021 | 1 | U | U02 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)U02.dat | 2 | 32 | 3272.9 | 24365 | 0.0097
29/09/2021 | 1 | CRM | CRM-189-0099-2 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-189-0099-2.dat | 2 | 33.494 | 3038.3 | 22619 | 0.148
29/09/2021 | 1 | CRM | CRM-189-0099-3 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-189-0099-3.dat | 2 | 33.494 | 3035.5 | 22597 | 0.038
13/10/2021 | 2 | J | J08 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J08.dat | 7 | 32 | 167.27 | 1246 | 3.059
13/10/2021 | 2 | J | J09 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J09.dat | 7 | 32 | 176.5 | 1315 | 0.581
13/10/2021 | 2 | J | J10 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J10.dat | 7 | 32 | 3153 | 23482 | 0.131 | Restarted laptop + Licor before this one
13/10/2021 | 2 | J | J11 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J1.dat | 7 | 32 | 3168.1 | 23592 | 0.144 | Restarted software before this one
13/10/2021 | 2 | J | J12 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J12.dat | 7 | 32 | 3179.3 | 23675 | 0.015
13/10/2021 | 2 | J | J13 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J13.dat | 7 | 32 | 3186.2 | 23727 | 0.026
13/10/2021 | 2 | J | J14 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J14.dat | 7 | 32 | 3189.4 | 23750 | 0.043
13/10/2021 | 2 | J | J15 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J15.dat | 7 | 32 | 3189 | 23746 | 0.005
13/10/2021 | 2 | J | J16 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J16.dat | 7 | 32 | 3191.6 | 23765 | 0.054
13/10/2021 | 2 | J | J17 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J17.dat | 7 | 32 | 3189.7 | 23751 | 0.114
13/10/2021 | 2 | CRM | CRM-189-0102-1000 | 360 | 75 | 120 | 10 | 1000 | 1000 | 150 | 150 | 15 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-189-0102-1000.dat | 2 | 33.494 | 2947.4 | 18289 | 0.073 | replicate #2 - subsampled from 500 mL bottle
13/10/2021 | 2 | CRM | CRM-189-0102-1200 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-189-0102-1200.dat | 2 | 33.494 | 2960.8 | 22047 | 0.062 | replicate #4 - subsampled from 500 mL bottle
13/10/2021 | 2 | CRM | CRM-189-0102-1400 | 360 | 75 | 120 | 10 | 1400 | 1400 | 150 | 150 | 21 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-189-0102-1400.dat | 2 | 33.494 | 2974.8 | 25844 | 0.089 | replicate #9 - subsampled from 500 mL bottle
13/10/2021 | 2 | R | R38 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)R38.dat | 2 | 32 | 3196.7 | 23803 | 0.024
13/10/2021 | 2 | R | R10 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)R10.dat | 2 | 32 | 3201.6 | 23839 | 0.08
13/10/2021 | 2 | U | U45 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)U45.dat | 2 | 32 | 3200.7 | 23832 | 0.028
13/10/2021 | 2 | R | R31 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)R31.dat | 2 | 32 | 3200.2 | 23829 | 0.032
13/10/2021 | 2 | U | U35 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)U35.dat | 2 | 32 | 3193.8 | 23781 | 0.119
13/10/2021 | 2 | U | U25 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)U25.dat | 2 | 32 | 3196 | 23798 | 0.104
13/10/2021 | 2 | R | R21 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)R21.dat | 2 | 32 | 3195.8 | 23795 | 0.092
13/10/2021 | 2 | U | U01 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)U01.dat | 2 | 32 | 3193 | 23774 | 0.087
13/10/2021 | 2 | U | U32 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)U32.dat | 2 | 32 | 3194.7 | 23787 | 0.107
13/10/2021 | 2 | R | R46 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)R46.dat | 2 | 32 | 3197.6 | 23808 | 0.051
13/10/2021 | 2 | CRM | CRM-189-0102-1200-2 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-189-0102-1200-2.dat | 2 | 33.494 | 2060.2 | 22041 | 0.145 | replicate #3 - subsampled from 500 mL bottle
27/10/2021 | 3 | J | J18 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J18.dat | 7 | 32 | 2147.9 | 23520 | 0.444
27/10/2021 | 3 | J | J19 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J19.dat | 7 | 32 | 2109.4 | 23195 | 0.117
27/10/2021 | 3 | J | J20 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J20.dat | 7 | 32 | 2109.3 | 23193 | 0.101
27/10/2021 | 3 | J | J21 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J21.dat | 7 | 32 | 2107.6 | 23179 | 0.261
27/10/2021 | 3 | J | J22 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J22.dat | 7 | 32 | 2102.9 | 23123 | 0.05
27/10/2021 | 3 | J | J23 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J23.dat | 7 | 32 | 2102.9 | 23123 | 0.127
27/10/2021 | 3 | J | J24 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J24.dat | 7 | 32 | 2104.6 | 23142 | 0.048
27/10/2021 | 3 | J | J25 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J25.dat | 7 | 32 | 2114.1 | 23246 | 0.2
27/10/2021 | 3 | J | J26 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J26.dat | 7 | 32 | 2122.5 | 23339 | 0.056
27/10/2021 | 3 | J | J27 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J27.dat | 7 | 32 | 2236.3 | 24590 | 0.038
27/10/2021 | 3 | J | J28 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J28.dat | 7 | 32 | 2233.3 | 24557 | 0.012
27/10/2021 | 3 | J | J29 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J29.dat | 7 | 32 | 2244.8 | 25584 | 0.042
27/10/2021 | 3 | J | J30 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J30.dat | 7 | 32 | 2232.3 | 24546 | 0.014
27/10/2021 | 3 | J | J31 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J31.dat | 7 | 32 | 2232.3 | 24546 | 0.037
27/10/2021 | 3 | CRM | CRM-189-0159-1000 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-189-0159-1000.dat | 2 | 33.494 | 2056.9 | 22617 | 0.042 | THIS SHOULD BE CALLED 1200
27/10/2021 | 3 | CRM | CRM-189-0159-1000-2 | 360 | 75 | 120 | 10 | 1000 | 1000 | 150 | 150 | 15 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-189-0159-1000-2.dat | 2 | 33.494 | 2045.1 | 18740 | 0.064 | THIS IS THE ONE THAT'S ACTUALLY 1000
27/10/2021 | 3 | CRM | CRM-189-0159-1400-1 | 360 | 75 | 120 | 10 | 1400 | 1400 | 150 | 150 | 21 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-189-0159-1400-1.dat | 2 | 33.494 | 2055.9 | 26374 | 0.023
27/10/2021 | 3 | R | R19 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)R19.dat | 2 | 32 | 2114.4 | 24349 | 0.128
27/10/2021 | 3 | J | J32 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J32.dat | 7 | 32 | -9999 | -9999 | -9999 | ABORTED - Kloehn not responding // restarted entire system because not recording analysis temperature
27/10/2021 | 3 | J | J33 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J33.dat | 7 | 32 | 2190.9 | 23991 | 0.039 | Kloehn responding but getting weird areas
27/10/2021 | 3 | J | J34 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J34.dat | 7 | 32 | 2188.3 | 23963 | 0.091 | Peltier ~-6.5
27/10/2021 | 3 | J | J35 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J35.dat | 7 | 32 | 2186.3 | 23941 | 0.078 | Peltier ~-6.5
27/10/2021 | 3 | J | J36 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J36.dat | 7 | 32 | 2180.6 | 23879 | 0.051 | Peltier ~-6.5
27/10/2021 | 3 | J | J37 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J37.dat | 7 | 32 | 2177.5 | 23845 | 0.058 | Peltier ~-6.5
29/10/2021 | 4 | J | J38 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J38.dat | 7 | 32 | -9999 | -9999 | -9999 | Aborted - restarting laptop - sound of 'disconnecting'
29/10/2021 | 4 | J | J39 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J39.dat | 7 | 32 | -9999 | -9999 | -9999 | Aborted - RESET ZERO before sample - Kloehn not reset
29/10/2021 | 4 | J | J40 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J40.dat | 7 | 32 | 2007.7 | 21991 | 0.184 | RESET ZERO before sample - bad integration curves, CO2 content too (-)13
29/10/2021 | 4 | J | J41 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J41.dat | 7 | 32 | 2006.9 | 21981 | 0.027 | bad integration curves
29/10/2021 | 4 | J | J42 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J42.dat | 7 | 32 | 2005.9 | 21970 | 0.144 | BEFORE SAMPLE: restarted laptop + RESET ZERO - still bad integration curves
29/10/2021 | 4 | J | J43 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J43.dat | 7 | 32 | 2241.9 | 24533 | 0.178 | RESET ZERO before sample - good integration curves
29/10/2021 | 4 | J | J44 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J44.dat | 7 | 32 | 2240.3 | 24534 | 0.22
29/10/2021 | 4 | J | J45 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J45.dat | 7 | 32 | 2240.4 | 24536 | 0.205
29/10/2021 | 4 | CRM | CRM-189-0159-1000-3 | 360 | 75 | 120 | 10 | 1000 | 1000 | 150 | 150 | 15 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-189-0159-1000-3.dat | 2 | 33.494 | 2063.6 | 18832 | 0.091 | subsample #3/7
29/10/2021 | 4 | CRM | CRM-189-0159-1200-1 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-189-0159-1200-1.dat | 2 | 33.494 | 2071.4 | 22684 | 0.24 | subsample #6/7
29/10/2021 | 4 | CRM | CRM-189-0159-1400-2 | 360 | 75 | 120 | 10 | 1400 | 1400 | 150 | 150 | 21 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-189-0159-1400-2.dat | 2 | 33.494 | 2075.3 | 26513 | 0.023 | subsample #7/7
29/10/2021 | 4 | R | R42 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)R42.dat | 2 | 32 | 2231.6 | 24437 | 0.045
29/10/2021 | 4 | R | R23 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)R42.dat | 2 | 32 | 2236.5 | 24491 | 0.186
29/10/2021 | 4 | U | U09 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)U09.dat | 2 | 32 | 2234.8 | 24472 | 0.062
29/10/2021 | 4 | U | U22 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)U22.dat | 2 | 32 | 2234.5 | 24470 | 0.127
29/10/2021 | 4 | U | U44 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)U44.dat | 2 | 32 | 2234.6 | 24471 | 0.018
29/10/2021 | 4 | CRM | CRM-189-0468-10 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-189-0468-10.dat | 2 | 33.494 | 2075.2 | 22725 | 0.213 | test subsample from 40 subsamples from 20/10/2021
29/10/2021 | 4 | R | R24 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)R24.dat | 2 | 32 | 2232.7 | 24450 | 0.027
29/10/2021 | 4 | R | R28 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)R28.dat | 2 | 32 | 2234.8 | 24473 | 0.073
29/10/2021 | 4 | U | U33 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)U33.dat | 2 | 32 | 2237.7 | 24504 | 0.183
29/10/2021 | 4 | U | U12 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)U12.dat | 2 | 32 | 2229.8 | 24417 | 0.201
29/10/2021 | 4 | CRM | CRM-189-0159-1200-2 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-189-0159-1200-2.dat | 2 | 33.494 | 2073.5 | 22706 | 0.247 | subsample #4/7
29/10/2021 | 4 | CRM | CRM-189-0468-09 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-189-0468-09.dat | 2 | 33.494 | 2069.5 | 22661 | 0.066 | test subsample from 40 subsamples from 20/10/2021
24/11/2021 | 5 | J | J46 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J46.dat | 7 | 32 | 2333.2 | 25558 | 0.143 | Pelter too cold and stays off while on in software
24/11/2021 | 5 | J | J47 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J47.dat | 7 | 32 | -9999 | -9999 | -9999 | Kloehn no working - RESTARTED SYSTEM + RESET ZERO x 1
24/11/2021 | 5 | J | J48 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J48.dat | 7 | 32 | -9999 | -9999 | -9999 | Kloehn no working - RESTARTED SYSTEM + RESET ZERO x 1
24/11/2021 | 5 | J | J49 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J49.dat | 7 | 32 | 2225.3 | 24375 | 0.316 | bad integration curves
24/11/2021 | 5 | J | J50 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J50.dat | 7 | 32 | 2158.1 | 23637 | 0.263 | Restarted software before this one
24/11/2021 | 5 | J | J51 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J51.dat | 7 | 32 | 2157.1 | 23626 | 0.151 | bad integration curves
24/11/2021 | 5 | J | J52 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J52.dat | 7 | 32 | 2333.1 | 25553 | 0.102 | GOOD - Reset zero before this one
24/11/2021 | 5 | J | J53 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J53.dat | 7 | 32 | 2335.1 | 25574 | 0.105 | GOOD
24/11/2021 | 5 | J | J54 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J54.dat | 7 | 32 | 2329.7 | 25514 | 0.171 | GOOD
24/11/2021 | 5 | J | J55 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J55.dat | 7 | 32 | -9999 | -9999 | -9999
26/11/2021 | 6 | J | J61 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J61.dat | 7 | 32 | -9999 | -9999 | -9999 | ABORTED - reset zero, Kloehn not working
26/11/2021 | 6 | J | J62 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J62.dat | 7 | 32 | -9999 | -9999 | -9999 | ABORTED - Peltier not coming on and pCO2 ~50
26/11/2021 | 6 | J | J63 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J63.dat | 7 | 32 | 2176.7 | 23844 | 0.147 | Restarted software - reset zero --> No integration curves for first 2
26/11/2021 | 6 | J | J64 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J64.dat | 7 | 32 | 2177.4 | 23851 | 0.076 | bad integration curves
26/11/2021 | 6 | J | J65 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J65.dat | 7 | 32 | 2183.8 | 23918 | 0.041 | bad integration curves
26/11/2021 | 6 | J | J66 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J66.dat | 7 | 32 | 2177.3 | 23847 | 0.067 | bad integration curves (negative pco2)
26/11/2021 | 6 | J | J67 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J67.dat | 7 | 32 | 2339 | 25767 | 0.042 | Reset zero - better curves
26/11/2021 | 6 | J | J68 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J68.dat | 7 | 32 | 2335.6 | 25580 | 0.119 | Reset zero - good curves
26/11/2021 | 6 | J | J69 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J69.dat | 7 | 32 | 2338.1 | 25606 | 0.157 | NO reset zero - good curves
26/11/2021 | 6 | J | J70 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J70.dat | 7 | 32 | 2329 | 25505 | 0.151 | NO reset zero - good curves
26/11/2021 | 6 | J | J71 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J71.dat | 7 | 32 | 2333.3 | 25552 | 0.136 | NO reset zero - good curves
26/11/2021 | 6 | J | J72 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J72.dat | 7 | 32 | 2334.3 | 25562 | 0.095 | NO reset zero - good curves
26/11/2021 | 6 | CRM | CRM-195-0078-1000 | 360 | 75 | 120 | 10 | 1000 | 1000 | 150 | 150 | 15 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-195-0078-1000.dat | 7 | 33.485 | 2074.9 | 18935 | 0.149
26/11/2021 | 6 | CRM | CRM-195-0078-1200 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-195-0078-1200.dat | 2 | 33.485 | 2081.9 | 22798 | 0.074
26/11/2021 | 6 | CRM | CRM-195-0078-1400 | 360 | 75 | 120 | 10 | 1400 | 1400 | 150 | 150 | 21 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-195-0078-1400.dat | 2 | 33.485 | 2087.3 | 26666 | 0.12
26/11/2021 | 6 | R | R50 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)R50.dat | 2 | 32 | 2228.1 | 24398 | 0.194
26/11/2021 | 6 | U | U13 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)U13.dat | 2 | 32 | 2226 | 24376 | 0.197
26/11/2021 | 6 | R | R27 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)R27.dat | 2 | 32 | 2228.4 | 24401 | 0.157
26/11/2021 | 6 | R | R07 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)R07.dat | 2 | 32 | 2225 | 24364 | 0.272
26/11/2021 | 6 | U | U19 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)U19.dat | 2 | 32 | 2227.6 | 24393 | 0.149
26/11/2021 | 6 | CRM | CRM-189-0468-07 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-189-0468-07.dat | 2 | 33.494 | 2072.7 | 22696 | 0.222
26/11/2021 | 6 | R | R05 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)R05.dat | 2 | 32 | 2224.9 | 24362 | 0.13
26/11/2021 | 6 | U | U34 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)U34.dat | 2 | 32 | 2228 | 24396 | 0.127
26/11/2021 | 6 | U | U20 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)U20.dat | 2 | 32 | 2226.5 | 24379 | 0.154
26/11/2021 | 6 | R | R30 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)R30.dat | 2 | 32 | 2231 | 24429 | 0.069
26/11/2021 | 6 | U | U37 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)U37.dat | 2 | 32 | 2226.9 | 24384 | 0.141
26/11/2021 | 6 | CRM | CRM-195-0078-1200-2 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-195-0078-1200-2.dat | 2 | 33.485 | 2076.6 |  |  | Forgot to write down values
26/11/2021 | 6 | CRM | CRM-195-0078-1200-3 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-195-0078-1200-3.dat | 2 | 33.485 | 2080 | 22776 | 0.136
20/01/2022 | 7 | J | J73 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J78.dat | 7 | 32 | -9999 | -9999 | -9999 | Aborted - Kloehn not pulling sample
20/01/2022 | 7 | J | J74 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J78.dat | 7 | 32 | -9999 | -9999 | -9999 | Aborted - Kloehn not pulling sample
20/01/2022 | 7 | J | J75 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J78.dat | 7 | 32 | -9999 | -9999 | -9999 | Restarted software prior to sample - Aborted - Kloehn ok but Peltier staying on while off in software
20/01/2022 | 7 | J | J76 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J78.dat | 7 | 32 | -9999 | -9999 | -9999 | Restarted system prior to sample - Aborted - Kloehn not ok, Peltier ok
20/01/2022 | 7 | J | J77 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J78.dat | 7 | 32 | -9999 | -9999 | -9999 | Restarted system prior to sample - Aborted - Kloehn not ok (no sound), Peltier ok
20/01/2022 | 7 | J | J78 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J78.dat | 7 | 32 | -9999 | -9999 | -9999 | Kloehn ok, Peltier ok - Reset zero prior to sample - LiCor not recording? Nothing on graph and CO2 ~ --500
20/01/2022 | 7 | J | J79 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J79.dat | 7 | 32 | 2362.1 | 25874 | 0.041 | Reset zero prior to sample - all above ok
20/01/2022 | 7 | J | J80 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J80.dat | 7 | 32 | 2361.4 | 25865 | 0.153 | Curves not amazing (especially 3rd one)
20/01/2022 | 7 | J | J81 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J81.dat | 7 | 32 | 2364 | 25894 | 0.183 | Curves not amazing (especially 3rd one)
20/01/2022 | 7 | J | J82 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J82.dat | 7 | 32 | 2358 | 25827 | 0.129 | Reset zero prior to sample - good curves
20/01/2022 | 7 | J | J83 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J83.dat | 7 | 32 | 2359.3 | 25840 | 0.248 | NO reset zero - good curves
20/01/2022 | 7 | J | J84 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J84.dat | 7 | 32 | 2359.3 | 25840 | 0.172 | NO reset zero - good curves
20/01/2022 | 7 | CRM | CRM-195-0114-1000 | 360 | 75 | 120 | 10 | 1000 | 1000 | 150 | 150 | 15 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-195-0114-1000.dat | 2 | 33.485 | 2075.7 | 18945 | 0.164 | NO reset zero - good curves
20/01/2022 | 7 | CRM | CRM-195-0114-1200 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-195-0114-1200.dat | 2 | 33.485 | 2081.5 | 22796 | 0.035 | Beautiful curves
20/01/2022 | 7 | CRM | CRM-195-0114-1400 | 360 | 75 | 120 | 10 | 1400 | 1400 | 150 | 150 | 21 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-195-0114-1400.dat | 2 | 33.485 | 2083.1 | 26616 | 0.135 | Not so great 4th curve
20/01/2022 | 7 | R | R49 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)R49.dat | 2 | 32 | 2225.5 | 24372 | 0.135
20/01/2022 | 7 | U | U50 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)U50.dat | 2 | 32 | 2231.1 | 24433 | 0.166
20/01/2022 | 7 | U | U07 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)U07.dat | 2 | 32 | 2226.5 | 24383 | 0.204
20/01/2022 | 7 | R | R18 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)R18.dat | 2 | 32 | 2225.8 | 24375 | 0.138
20/01/2022 | 7 | R | R25 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)R25.dat | 2 | 32 | 2225 | 24366 | 0.17
20/01/2022 | 7 | U | U38 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)U38.dat | 2 | 32 | 2228.2 | 24401 | 0.05
20/01/2022 | 7 | CRM | CRM-189-0468-17 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-189-0468-17.dat | 2 | 33.494 | 2068.4 | 22651 | 0.15
20/01/2022 | 7 | U | U11 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)U11.dat | 2 | 32 | 2228.6 | 24404 | 0.17
20/01/2022 | 7 | R | R34 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)R34.dat | 2 | 32 | 2228 | 24398 | 0.067
20/01/2022 | 7 | R | R35 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)R35.dat | 2 | 32 | 2227.3 | 24391 | 0.026
20/01/2022 | 7 | U | U21 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)U21.dat | 2 | 32 | 2227.5 | 24392 | 0.231
20/01/2022 | 7 | CRM | CRM-195-0078-25 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-195-0078-25.dat | 2 | 33.485 | 2078.1 | 22756 | 0.29
20/01/2022 | 7 | CRM | CRM-195-0114-1200-2 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-195-0114-1200-2.dat | 2 | 33.485 | 2081.5 | 22794 | 0.215
2022-02-16 |  | J | J85 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J85.dat | 7 | 32 | 2189.4 | 23985 | 0.225
2022-02-16 |  | J | J86 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J86.dat | 7 | 32 | 2248.1 | 24720 | 0.068
2022-02-16 |  | J | J87 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J87.dat | 7 | 32 | 2249.6 | 24736 | 0.053
2022-02-16 |  | J | J88 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J88.dat | 7 | 32 | 2247.0 | 24707 | 0.046
2022-02-16 |  | J | J89 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J89.dat | 7 | 32 | 2246.0 | 24696 | 0.099
2022-02-16 |  | J | J90 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J90.dat | 7 | 32 | 2242.8 | 24661 | 0.028
2022-02-16 |  | J | J91 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J91.dat | 7 | 32 | 2121.7 | 23139 | 5920
2022-02-16 |  | J | J92 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J92.dat | 7 | 32 | 2176.9 | 23842 | 0.115
2022-02-16 |  | J | J93 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J93.dat | 7 | 32 | 2266.1 | 24819 | 0.286
2022-02-16 |  | J | J94 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J94.dat | 7 | 32 | 2263.9 | 24795 | 0.199
2022-02-16 |  | J | J95 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J95.dat | 7 | 32 | 2266.4 | 24822 | 0.088
2022-02-16 |  | J | J96 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J96.dat | 7 | 32 | 2273.0 | 24892 | 0.315
2022-02-16 |  | J | J97 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J97.dat | 7 | 32 | 2268.7 | 24844 | 0.172
2022-02-16 |  | J | J98 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J98.dat | 7 | 32 | 2265.4 | 24808 | 0.296
2022-02-16 |  | J | J99 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J99.dat | 7 | 32 | 2269.4 | 24852 | 0.222
2022-02-16 |  | J | J100 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J100.dat | 7 | 32 | 2265.4 | 24809 | 0.235
2022-02-16 |  | J | J101 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J101.dat | 7 | 32 | 2265.9 | 24814 | 0.202
2022-02-17 |  | J | J102 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J102.dat | 7 | 32 | 2184.4 | 23939 | 0.342
2022-02-17 |  | J | J103 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J103.dat | 7 | 32 | 2192.9 | 24026 | 0.223
2022-02-17 |  | J | J104 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J104.dat | 7 | 32 | 2194.1 | 24038 | 0.271
2022-02-17 |  | J | J105 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J105.dat | 7 | 32 | 2217.9 | 24292 | 0.076
2022-02-17 |  | J | J106 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J106.dat | 7 | 32 | 2221.7 | 24334 | 0.303
2022-02-17 |  | J | J107 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J107.dat | 7 | 32 | 2263.6 | 24793 | 0.278
2022-02-17 |  | J | J108 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J108.dat | 7 | 32 | 2260.8 | 24762 | 0.179
2022-02-17 |  | J | J109 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)J109.dat | 7 | 32 | 2260.6 | 24759 | 0.124
2022-02-17 |  | CRM | CRM-189-0408-1200-10 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-189-0408-1200-10.dat | 2 | 33.494 | 2077.2 | 22749 | 0.268
2022-02-17 |  | CRM | CRM-189-0408-1000-2 | 360 | 75 | 120 | 10 | 1000 | 1000 | 150 | 150 | 15 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-189-0408-1000-2.dat | 2 | 33.494 | 2069.3 | 18885 | 0.064
2022-02-17 |  | CRM | CRM-189-0408-1200-4 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-189-0408-1200-4.dat | 2 | 33.494 | 2072.8 | 22701 | 0.067
2022-02-17 |  | CRM | CRM-189-0408-1400-6 | 360 | 75 | 120 | 10 | 1400 | 1400 | 150 | 150 | 21 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-189-0408-1400-6.dat | 2 | 33.494 | 2076.1 | 26527 | 0.235
2022-02-17 |  | U | U16 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)U16.dat | 2 | 32 | 2237.2 | 24500 | 0.222
2022-02-17 |  | R | R9 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)R9.dat | 2 | 32 | 2240.0 | 24531 | 0.233
2022-02-17 |  | R | R16 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)R16.dat | 2 | 32 | 2237.7 | 24505 | 0.220
2022-02-17 |  | U | U18 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)U18.dat | 2 | 32 | 2237.4 | 24502 | 0.277
2022-02-17 |  | U | U31 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)U31.dat | 2 | 32 | 2238.5 | 24514 | 0.222
2022-02-17 |  | CRM | CRM-189-0408-1200-7 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-189-0408-1200-7.dat | 2 | 33.494 | 2068.8 | 22655 | 0.516
2022-02-17 |  | CRM | CRM-189-0408-1200-9 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-189-0408-1200-9.dat | 2 | 33.494 | 2074.3 | 22715 | 0.295
2022-02-17 |  | U | U23 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)U23.dat | 2 | 32 | 2235.1 | 24475 | 0.050
2022-02-17 |  | R | R48 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)R48.dat | 2 | 32 | 2233.1 | 24454 | 0.087
2022-02-17 |  | R | R14 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)R14.dat | 2 | 32 | 2235.0 | 24475 | 0.264
2022-02-17 |  | U | U15 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)U15.dat | 2 | 32 | 2230.5 | 24428 | 0.259
2022-02-17 |  | R | R32 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)R32.dat | 2 | 32 | 2227.7 | 24397 | 0.148
2022-02-17 |  | CRM | CRM-189-0408-1200-3 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-189-0408-1200-3.dat | 2 | 33.494 | 2069.5 | 22665 | 0.175
2022-02-17 |  | CRM | CRM-189-0408-1200-8 | 360 | 75 | 120 | 10 | 1200 | 1200 | 150 | 150 | 18 | 3 | 2 | all_to_stripped | no | LD_storage_test.dbs | 0-0 0 (0)CRM-189-0408-1200-8.dat | 2 | 33.494 | 2067.4 | 22643 | 0.018